## Slide 1: Application of Micro-MEGAS on measuring X-ray activity of 71Ge
Kangkang Zhao
State Key Laboratory of Particle Detection and Electronics
University of Science and Technology of China
2021, 10. 22 XiAn, China

## Slide 2: Outline
NTD-Ge thermistor
Ge-71 thermal neutron fluence monitor
Micro-MEGAS: X-ray detection
Geant4 simulation
Results
Summary

## Slide 3: NTD-Ge thermistor
Neutron Transmutation Doped semi-conductor (Germanium)
The net carrier concentration is
For a natural Ge, the neutron doping results in p-type due to its isotopic abundance and thermal neutron capture cross section
Isotope | Abundance (%) | Cross-section (bar) | Production (half-life) | Doped product | Dopant type
70Ge | 20.55 | 3.05(13) | 71Ge (11.2 d) | 71Ga (EC) | Acceptor p-type
72Ge | 27.37 | 0.89(8) | 73Ge (stable) | 73Ge | --
73Ge | 7.67 | 14.7(4) | 74Ge (stable) | 74Ge | --
74Ge | 36.74 | 0.36(4) | 75Ge (82.8 m) | 75As (- decay) | Donor n-type
76Ge | 7.67 | 0.055(2) | 77Ge (38.8, 11.3 h) | 77Se, 77As (-, -) | Donor n-type
Isotope | Abundance (%) | Cross-section (bar) | Production (half-life) | Doped product | Dopant type
70Ge | 20.55 | 3.05(13) | 71Ge (11.2 d) | 71Ga (EC) | Acceptor p-type
72Ge | 27.37 | 0.89(8) | 73Ge (stable) | 73Ge | --
73Ge | 7.67 | 14.7(4) | 74Ge (stable) | 74Ge | --
74Ge | 36.74 | 0.36(4) | 75Ge (82.8 m) |  | Donor n-type
76Ge | 7.67 | 0.055(2) | 77Ge (38.8, 11.3 h) |  | Donor n-type

## Slide 4: NTD-Ge thermistor
The resistivity of a NTD Ge is given by Mott’s law as,
In a Variable Hopping mechanism, the is predicted to be 0.5
neutron capture cross-section - homogenous doping
high temperature sensitivity, reproducibility, mass production, low specific heat, fast rise time and simple read out.
Ge intrinsic property
Doping level
0.02
0.04
0.06
0.08
0.10
0.12
102
103
104
105
106
Temperature (K)
Resistance (Ω)
3.0 mm
2.9 mm
0.9 mm
0.2 mm

## Slide 5: Neutron Irradiation of Ge
10Nine High-purity Ge, preparation and wrap
China Advanced Research reactor, CARR
2020.12.23, irradiation time 64 h 29 min
Thermal neutron percentage > 99%
No fluence monitor
Wafer dicing
HNO3:HF=7:3
etching and de-oxide
Wrap in quartz box
Sealed in Al can
10*10*1 mm3

## Slide 6: Thermal neutron fluence measurement
The thermal neutron fluence of HPGe received is related to the doping level of NTD-Ge, and the performance of thermistor
Using 70Ge(n, ) 71Ge reaction: 9.2 keV and 10.3 keV X-rays
Principle of measurement
The X-ray detection was carried out on MMD
71Ga

## Slide 7: NTD-Ge Radioactivity
Low-background HPGe gamma spectrometer
EC内轫致连续
Cr-51
Zn-65
Co-60
Co-60
Fe-59
Sb-124
，100% EC，Q=232.49keV

## Slide 8: Micro-MEGAS
9.2 keV and 10.3 keV X-rays activities measurement
Clear structure: geant4 simulation to obtain efficiency
Thermal bonding method manufactured MMD
Feng J, Zhang Z, Liu J, et al 2021, 989: 164958.
HPGe: 10*10*1mm3
Cu-collimator
PE film wrapping
No etching, no pollution
Easy to handle at lab
Easy to make new structures
Cheap
Φ0.5mm- Φ1mm spacers, ~1cm pitch
easy to clean, especially for large area
less than 1% spacer area

## Slide 9: Experimental setup
Micro-Megas: Gas mixture 95% Ne, 5% iC4H10
Determine the optimal work points of MMD
  Electronic transparency
  Fixed the mesh voltage, adjust drift voltage, choose the point that have maximum peak position
  Energy resolution vs gain
  Fixed the field ratio, adjust voltage
  value, choose the gain that near 5000
∅6.45mm Cu collimator
Anticoincidence system to reduce bkg

## Slide 10: 71Ge X-ray spectrum
Energy resolution: ~9.1% FWHM@9.2 keV
Fit:
  Signal: 5 gaussian characteristic X-ray
  Background: Argus + Chebyshev polynomials
Prominent features: 5.4keV, 6.4keV, 8.0keV, which are NOT seen in Fe-55 spectrum
~11.4% FWHM@5.9 keV
Notes: Penning effect

## Slide 11: Peak sources
8.0 keV peak is from the Cu collimator excitation
With Cu-collimator
Without Cu-collimator

## Slide 12: Half-life fit results
5.4 keV
6.4 keV
9.2 keV
10.3 keV
Energy (keV) | Half-life (day)
5.4 | 11.28
6.4 | 11.02
9.2 | 12.12
10.3 | 11.02
Half-life of 71Ge is 11.42 day, so the Prominent features: 5.4keV, 6.4keV are probably relative to 71Ge ( 9.2 keV , 10.3 keV )

## Slide 13: Mesh material excitation
Mesh material: stain-less mesh
Cr, Mn, Fe elements in stainless mesh are excited

## Slide 14: SDD measurement
Silicon Drift Detector
SDD Ge-71 spectrum
SDD
NTD-Ge
Mesh

## Slide 15: Geant4 simulation
71Ge nuclides, random and homogeneous and zero kinetic
energy placed in the NTD-Ge
by using the G4GPS
eﬀiciency calculation, =(3.5840.248) E-4
Cu collimator
Thick 5mm
D:6.45mm
PET window
15x15 cm2x330um
NTD-Ge
10x10x1 mm3
PE Film
8um
Gas chamber
15x15 cm2x1.1cm
95%Ne+5%iC4H10
Efficiency contribution
Source distribution
NTD-Ge Self absorption
Cu collimator acceptance
Window absorption
Interaction with gas

## Slide 16: Thermal neutron fluence results
Fluence results and the uncertainty budget
The results of MMD and SDD measured are consistent
Dominant uncertainty of MMD comes from the uncertainty of thickness drift gas chamber
误差来源 | 相对误差 (%) | 
 | MMD | SDD
截面 | 1.5385 | 1.5385
探测器效率 | 7.0335 | 2.1203
输入参数 | 2.9267 | 2.9268
半衰期 | 2.9847 | 2.8612
合计 | 8.3253 | 4.8683
preliminary
preliminary

## Slide 17: Summary
Using MMD to detect 9.2, 10.3 keV X-rays of 71Ge
Found the peaks from mesh material excitation, which are not seen by 55Fe source
Using Neon gas reached a good energy resolution of ~9.1% FWHM@9.2 keV
Geant4 simulation give the detection-efficiency of MMD
Firstly using Ge and MMD to measure thermal neutron fluence
Future on-line reactor experiment
Thank you!
Notes: 高流强 大面积

## Slide 18: Backup
MMD measured spectrum Fit results
With Cu-collimator
Without Cu-collimator

## Slide 19: SDD info
Be窗
多层准直
厚度100钨W+35铬Cr+15钛Ti+75铝Al
硅晶体
测试能区2-30 keV
能量分辨率150 eV@5.9keV (Fe55)
实验室测试实物图
9.22keV
10.26 keV
71GeEC衰变kX-ray能谱

## Slide 20: SDD geant4 simulation