## Slide 1: New Prototype MRPC with Diamond-Like-Carbon Coated Electrodes
Kaiyang Wang, Xinjian Wang, Yongjie Sun
State Key Laboratory of Particle Detection and Electronics
Department of Modern Physics, USTC

## Slide 2: Outline
Research background
Introduction to DLC technology
MRPC prototype and cosmic ray test
Problems and improvement methods
Summary

## Slide 3: Background
The Multi-gap Resistive Plate Chamber (MRPC) has been successfully used in Alice TOF, STAR eTOF, STAR MTD, BESIII eTOF
High time resolution <50ps
High detect efficiency >95%(MIPs)
Low price
The Compressed Baryonic Matter(CBM) experiment will use MRPC technology in TOF
A rate capability between 0.5kHz/cm² ~25kHz/cm²
Normal MRPC can’t meet this demand, high rate capability MRPC is needed

## Slide 4: rate capability of MRPC
Vappl: Voltage apply to MRPC
Φ: particle flux
ρ: resistivity of glass
d: thickness of glass
<Q>: average electric charge
Normal way
Decrease thickness of glass: ultra-thin float glass
Decrease resistivity of glass : low resistive glass (Tsinghua University)
New way
Coat a Diamond-Like-Carbon (DLC) layer all over the entire surface of glass to change path for the neutralization of the avalanche charges

## Slide 5: DLC: Diamond Like Carbon
A kind of “Metastable amorphous carbon material which contains both diamond-structure and graphite-structure”
Surface resistance stable and adjustable
Resistant to discharge and radiation
Chemical stability
Schematic of Magnetron Sputtering deposition
Target
Sample holder
State Key Laboratory of Solid Lubrication, Lanzhou Institute of Chemical Physics, Chinese Academy of Sciences
Manufactured by Magnetron Sputtering deposition

## Slide 6: U-I curve
12cm x 20cm
DLC glass sample
Made 2 DLC glass samples, surface resistivity are 9GΩ/□ and 90GΩ/□
Electronic transmission speed is faster in DLC glass

## Slide 7: MRPC prototype
3 kinds of MRPC: one normal MRPC
two DLC MRPC 9GΩ/□&90GΩ/□
2x3 gaps
gap width 0.23mm
8 readout strips
Strip: 200mm x 7mm
For DLC MRPC, all resistive plates use DLC glass
cosmic ray test system schematic
Signal picked up after FEE
90GΩ/□ DLC MRPC@±3600V(~104kV/cm)

## Slide 8: High voltage test
Leak current
Glass MRPC: <15nA@±4000V(~116kV/cm)
9GΩ/□ DLC MRPC: >100nA@±3200V(~93kV/cm)
90GΩ/□ DLC MRPC: <20nA @±4000V(After one month HV test, leak current increased)
high leak current
Problem1:
after a month

## Slide 9: Time resolution && efficiency
low detect efficiency <60%
Problem2:
Glass MRPC@±3600V
Efficiency ≈90%
Time resolution<80ps
Time resolution and Efficiency plateau
of 90GΩ/□ DLC MRPC

## Slide 10: Inspection on DLC MRPC
DLC layers have fallen out
Fallen out
Increase glass surface roughness
different electrode base
Ways to improve

## Slide 11: Erased Glass and Other Base
Erased glass: increase glass surface roughness
Normal float glass
Ra=0.37nm
NaOH erased glass
Ra=1.79nm
HF erased glass
Ra=3.58nm
Other base electrode: Kapton-PCB
DLC-Kapton-PCB
150GΩ/□
The problem of DLC shedding has not been solved so far
More test on other bases is ongoing

## Slide 12: CST simulation
Why low detect efficiency? DLC layers have shielded induce signal?
Build a simulation model by CST software
Change layer’s conductivity and thickness h to adjust surface resistivity
B1
B2
B3
T1
T2
T3
Source
Simulation model
Excitation signal
induced signal
induced signal varies surface resistivity

## Slide 13: Simulation result
DLC layer’ Shielding effect
Shielding effect is not the main reason for low detect efficiency!
Excitation | DLC glass | Glass
gap | 4.66nA | 4.72nA
gap | 4.09nA | 4.71nA
gap | 3.53nA | 4.55nA
All gaps | 12.28nA | 13.98nA
Excitation | DLC glass | Glass
 | 4.66nA | 4.72nA
 | 4.09nA | 4.71nA
 | 3.53nA | 4.55nA
All gaps | 12.28nA | 13.98nA
Induced signal
Excitation signal
Build more detail simulation model refer to our MRPC structure
For DLC layers, set surface resistivity as maximum value 1MΩ/□
Glass or DLC glass
Amplitude of induced signal
gap
gap
gap

## Slide 14: Investigation
two single stack MRPCs with 6 gas gaps
Different electrode and inner plates
MRPC1: DLC electrode + inner glass
MRPC2: Graphitic electrode + inner DLC(100G)
@±7200V(~104kV/cm)
Leak current: <40nA
Efficiency: 45.6%
Time resolution: ~100ps
@±7200V(~104kV/cm)
Leak current: <40nA
Efficiency: 86.7%
Time resolution: ~103ps
Detect efficiency is related to DLC electrode!
Need to test DLC coating + Graphitic glass
MRPC1:
MRPC2:

## Slide 15: Summary
Put up with a new way to enhance rate capability of MRPC
by DLC coating in electrode
Made different kinds of MRPC prototype, test by cosmic ray
Encountered two main problems : large leak current and low efficiency
Find the reason and put forward ways to improve
Future work
Strengthen adhesion between DLC coating and glass
Test graphitic glass + DLC coating
Rate capability research
Thank you!