## Slide 1
, , , ,
Study of DLC photocathode for PICOSEC detectors based on Micromegas
State Key Laboratory of Particle Detection and Electronics, University of Science and Technology of China
State Key Laboratory of Solid Lubrication, Lanzhou Institute of Chemical Physics, Chinese Academy of Science
State Key Laboratory of Transient Optics and Photonics , Xi'an Institute of Optics and Precision Mechanics, Chinese Academy of Sciences

## Slide 2
CEA Saclay(France):D. Desforge, I. Giomataris, T. Gustavsson, C. Guyot, F.J. Iguaz, M. Kebbiri, P. Legou, O. Maillard, T. Papaevangelou, M. Pomorski, P. Schwemling, L. Sohl.
CERN:J. Bortfeldt, F. Brunbauer, C. David, J. Frachi, M. Lupberger, H. Müller, E. Oliveri, F. Resnati, L. Ropelewski, T. Schneider, P. Thuiner, M. van Stenis, P. Thuiner, R. Veenhof, S. White2.
USTC (China):J. Liu, B. Qi, X. Wang, Z. Zhang, Y. Zhou.
AUTH (Greece):I. Manthos, V. Niaouris, K. Paraschou, D. Sampsonidis, S.E. Tzamarias.
NCSR:G. Fanourakis.
NTUA:Y. Tsipolitis.
LIP:M. Gallinaro.
HIP (FINLAND): F. García.
IGFAE: D. González-Díaz.
PICOSEC Collaboration

## Slide 3
Outline
Introduction
  PICOSEC detector concept
  Results with CsI photocathode
Diamond-Like Carbon (DLC) Photocathode
  Manufacture
  Quantum Efficiency test with laser
  Beam test
Conclusion and Future work

## Slide 4
Detector Concept
A Novel fast timing Micromegas detector: Cherenkov Light Detection
Cherenkov Radiator & Photocathode
Smaller drift gap
Higher electric field
Schematic of PICOSEC detector
Notes: Here is schematic of Picosec detector prototype. we reduce the thickness of the drift gap from a few millimeters to 200 μm, which will reduce the time jitter caused by the slow drift process. Then the Cherenkov radiator coated with a photocathode is coupled to the detector. The photoelectrons generation process is simultaneously. So the time jitter is very small. The smaller drift gap and higher electric field also have the positive effect to reduce the time jitter. The primary photoelectrons are pre-amplified in the drift gap. Next slide please.

## Slide 5
Started as an RD51 common fund project:
Fast Timing for High-Rate Environment: A Micormegas Solution
Collaboration
2014
2016
2015
2017
2018
Proposal
Submission
First Prototype and laser test
New Prototypes, laser tests and measurements with charged particles (CsI photocathode)
Resistive micromegas prototypes,
Multi-channel anode and larger area, photocathodes (CsI protection, Diamond, DLC),
Robust photocathode (DLC-based, pure metallic and nano diamond-based photocathodes), Double mesh structure, new detector prototypes… …
The crucial issue for a successful application of this concept is the robustness of photocathodes
Notes: Here is schematic of Picosec detector prototype. we reduce the thickness of the drift gap from a few millimeters to 200 μm, which will reduce the time jitter caused by the slow drift process. Then the Cherenkov radiator coated with a photocathode is coupled to the detector. The photoelectrons generation process is simultaneously. So the time jitter is very small. The smaller drift gap and higher electric field also have the positive effect to reduce the time jitter. The primary photoelectrons are pre-amplified in the drift gap. Next slide please.

## Slide 6
Results with CsI
Time resolution reaches: 28 ps
Mean number of photoelectrons per muon: 7.4
Detection efficiency for muons: 100%
Ion back-flow (IBF) ratio: 42.7%
USTC PICOSEC detector (18 nm CsI)
CsI aging problem
High IBF
mesh imprint on CsI layer
Sparks:
Crack on CsI layer
Cathode: -475V
Mesh: GND
Anode: +300V
Some alternative photocathodes need to be studied

## Slide 7
Diamond-Like Carbon (DLC)
Schematic of DLC deposition by magnetron sputtering technology
State Key Laboratory of Solid Lubrication, Lanzhou Institute of Chemical Physics, Chinese Academy of Science
More detailed in Lunlin Shang’s presentation
MgF2
DLC photocathode samples
3 mm MgF2 with DLC

## Slide 8
QE test with Laser
Relative QE of DLC with different thicknesses
Absolute QE is unpersuasive
Relative quantum efficiency of different samples are compared
Silicon detector (No DLC sample)
Incident photons flux
MCP detector (DLC sample)
Photoelectrons generation rate
Quantum efficiency test system in laboratory
State Key Laboratory of Transient Optics and Photonics , Xi'an Institute of Optics and Precision Mechanics, Chinese Academy of Sciences
The sample with thickness of 2.5 nm shows optimal performance
Notes: The QE was tested by UV laser in the State Key Laboratory of Transient Optics and Photonics in China. As shown in the left picture. the light emitted by the Deuterium lamp passes through monochromator and the wavelength of the out light ranges from 120 nm to 250 nm, The optical path calibration and detector system are placed in two vacuum chambers, which are separated by a 〖𝑀𝑔𝐹〗_2 crystal, which acts as an incident window for the second chamber. The testing sample (DLC photocathode) is placed behind the window, in front of the detector. Compare to the voltage applied to the detector, the negative voltage is applied in the DLC layer. Photoelectrons will be emitted from the surface of the DLC film when exposed to UV light. The photoelectrons drift to the MCP detector in the electric field, then we can get the events rate of photoelectrons through readout electronics. Then we remove the DLC photocathode and replace the MCP detector with a Silicon detector, which is calibrated by the National Institude of Standards and Technology. When the UV light is irradiated on the silicon detector, the flux of the incident photon can be known by the generated photocurrent. By dividing the test results of the MCP detector and the silicon detector, we can get the QE of the sample. The premise of using this way is that through the previous experiments, the intensity of the light source is stable enough during our test.
The results are shown in the right. Absolute QE is not very high. But we can compare the performance relatively. Because we have some preliminary results of 10 nm samples, so we use it as a reference. We found

## Slide 9
Beam Line
GEM Tracker
GEM Tracker
GEM Tracker
5mm hole veto scintillator
PICOSEC
PICOSEC
PICOSEC
DMM
10 cm * 10 cm scintillator
5 mm * 5 mm scintillator
Hamamatus MCP PMT
Scintillator
Beam test of PICOSEC Detector with DLC
Time reference
USTC PICOSEC detector
Active area: Φ1 cm
DLC Photocathode

## Slide 10
Data analysis
Blue : Sample waveform of PICOSEC detector generated by 150 GeV muon
Red : Related MCP PMT signal
Signal Arrival Time: time difference between PICOSEC and MCP PMT (Timing with CFD)
Single photoelectron calibration: UV lamp
Fit function: Polya distribution
Time Response
Npe: mean number of photoelectrons generated by Cherenkov light on the surface of DLC layer
f = 1/(1+θ): relative gain variance: mean

## Slide 11
Thickness of DLC film (nm) | Npe/per muon | Detection efficiency for muons
1 | Bad | Bad
2.5 | 3.7 | 97%
2.5 | 3.5 | 95%
2.5 | 3.4 | 94%
5 | 3.4 | 94%
7.5 | 2.2 | 70%
10 | 1.7 | 68%
5 nm Cr + 18 nm CsI | 7.4 | 100%
Npe results of different DLC samples
Beam tests results confirm the laser results
2.5 nm sample is currently the optimal one , with 3.7 Npe /per muon and 97% efficiency (the best results of several 2.5 nm samples)
These results are repeatable:

## Slide 12
Radiator MgF2 (mm) | Photocathode | Npe
3 | Cr + CsI(2 nm LiF layer) | <1
3 | Cr + CsI(2 nm MgF2 layer) | 3.55
3 | 20 nm Cr | 0.66
3 | 6 nm Al | 1.69
5 | 10 nm Cr | 2.15
5 | 10 nm Al | 2.20
3 | 2.5 nm DLC | 3.7
3 | 5 nm DLC | 3.4
Npe results of different photocathodes
Different photocathodes have been studied
DLC is promising compared with others

## Slide 13
Time resolution with DLC photocathode
Time resolution reaches 42 ps with 2.5 nm DLC photocathode
Higher QE of photocathode, better time resolution of detector will be reached

## Slide 14
Robustness of DLC photocathode
Evidence of DLC’s stability
Stored in atmospheric environment for a few months, without any extra protection
Worked stable during beam period
No performance degradation by muons for a few days
Pions irradiation
2.5 nm DLC photocathode was placed in a Resistive PICOSEC detector (-525/+275)
3.5 h: pions trigger number , 0.5 mC / 0.8 cm2 on DLC layer
(no shower is considered, actually charge should be higher)
 | Npe/per muon | Detection efficiency
Before pions | 3.5 | 95%
After pions | 3 | 94%

## Slide 15
PICOSEC detector based on
Double Mesh Micromegas
Transmission Type
Reflection Type
Low IBF (180 ps @2 ‰ IBF ratio CsI photocathode)
Photocathode blind to IBF
Higher QE
https://indico.ihep.ac.cn/event/7352/session/10/contribution/460
CsI
Need more optimization

## Slide 16
Conclusion
Laser test: 2.5 nm sample shows best performance
Beam test: 2.5 nm sample is also the best one
3.7 Npe / per muon 97 % detection efficiency
time resolution reaches 42 ps
After pions irradiation, QE performance shows less decrease, more resistant to irradiation than CsI
For DLC’s resistive character, it acts as a resistive cathode that restrains sparks
DMM preliminary results:
    Transmission type: low IBF ratio(180 ps @2 ‰)
    Reflection type: no good results

## Slide 17
Future Work
Find a suitable thickness of DLC layer which is robust enough, as well as suitable QE
Optimize the process of DLC deposition:
doping, heating… …
Study Secondary Electron Emission of DLC
Improve the QE test system
Aging research in lab with laser (213 nm):
    Figure out the impact of IBF
Continue working in DMM
… …