## Slide 1: MRPC 飞行时间谱仪技术研究进展
Yi Wang
Department of Engineering Physics
Tsinghua University

## Slide 2: Outline
Introduction of MRPC -TOF system
Future requirement for MRPC-TOF:
-- High rate >20kHz/cm2 (FAIR-CBM)
-- High time resolution <20ps (Jlab-SoLID)
-- Sealed MRPC to reduce gas consumption
Conceptual design MRPC for STCF endcap
Summary

## Slide 3: MRPC introduction
Application in nuclear physics experiments
Application in industry（Muon tomography)
Application in medicine（TOF-PET）
MRPC application：
MRPC structure
Main performance:
Excellent time： ->20 ps
Efficiency：>98%
Spatial resolution: ~mm
Cost effective
Long time stability！
Working mechanism:
Electrode: glass, ~1012.cm
HV layer: graphite, ~M /
Narrow gap: 100-300 m
High electric field: >105 V/cm
Working gas should be very electro-negative:
90%R134a+ 5%Iso-butane+5%SF6

## Slide 4: MRPC in ToF system
MRPC has been widely adopted to construct the Time of Flight (TOF) systems in HEP experiments.
 | ALICE | STAR | FOPI | BESIII | CBM | SoLID
Active area per detector (cm) | 120 x 13 | 22 x 8.4 | 90 x 4.6 | 0.5x(9.2+14.8)x32.8 | 33 x 27.6 | --
Total active area (m2) | 141 | 50 | 5 | 1.33 | 120 | 10
Pad size (cm) | 3.7 x 2.5 | 6.3 x 3.1 | 90 x 0.3 | (9.1~14.1) x 2.4 | 27 x 1.0 | (16~28) x 2.5
Gap×thickness(mm) | 10 x 0.25 | 6 x 0.22 | 6 x 0.3 | 12 x 0.22 | 10 x 0.25 | 32 x 0.128
Gas mixtures (C2H2F4/ C4H10/SF6) | 90/5/5 | 95/5/0 | 85/5/10 | 90/5/5 | 90/5/5 | 90/5/5
Operating field (kV/cm) | 96 | 107 | 110 | 109 | 110 | 140
Efficiency | 99.9% | 95-97% | 97±3% | 99% | 97% | 98%
Time resolution(ps) | 40 | 60 | 73±5 | 60 | 60 | 20 ps
Max rate (Hz/cm2) | 50 | 10 | 50 | 50 | 30k | 20k
higher counting rate and time precision.

## Slide 5: RHIC-STAR TOF system
STAR TOF-Tray

## Slide 6: PID and physics success
TOF
TOF PID:
 /k ~1.6 GeV/c,
(,k)/p ~ 3.0 GeV/c
Nature 473 (2011) 353

## Slide 7
Primary Beams
1012/s; 1.5 GeV/u; 238U28+
1010/s 238U73+ up to 35 GeV/u
3x1013/s 30 GeV protons
The „FAIR 2028“ science program will include:
  CBM at the new cave with SIS100 beams, and HADES at SIS18
  APPA experiments at the low-energy rings, at SIS100, at the caves at SIS18 and UNILAC with and at PHELIX and a limited set of experiments which could be hosted at all the caves served by SIS100
FAIR 2028

## Slide 8
8
8
CBM-ToF Requirements
Full system time resolution sT ~ 80 ps
Efficiency > 95 %
Rate capability  25 kHz/cm2
Polar angular range 2.5°– 25°
Occupancy < 5 %
Low power electronics
(~100.000 channels)
Free streaming data acquisition
The structure of CBM-TOF wall

## Slide 9: Higher counting rate: low-resistive glass
low-resistive glass control the voltage drop (efficiency loss) when incident flux goes up by Decrease the resistivity of the electrodes
Low-resistive glass has been operating at FAIR-Phase 0 programs like STAR-eTOF and mCBM
It has been applied in CBM-ToF Wall
Beam test result:
93%, 80ps ,70kHz/cm2
Specifications of low resistive glass
Notes: 在介绍为

## Slide 10: Higher counting rate : gas pollution
Constant rate condition of 1kHz/cm2
The gas pollution in the gas gap caused the increase of current!
Gas tube
Inlet/outlet
Fishline hook
3D printed sealing frame with Good strength, insulation and radiation persistency
Sealing MRPC speed up gas exchange and can reduce to 1sccm/m2 flush rate at cosmic experiment!

## Slide 11: Excellent performance of sealed MRPC
Sealed MRPC promoted gas exchange
Excellent current behavior under high rate irradiation:
Stable current with constant rate condition.
Fast decay of dark current since when X-ray is off
Decrease the wait time of gas purging:
Reach the working HV in 2h since flowing the gas

## Slide 12: High counting rate (>7kHz/cm2): creepage on fishing line
The creepage on the fishing line caused the high current and deposition!
35 um
fishline contact
Chart: Relaxation time [h]
| 1 | 1 | 2 | 10 | 60 | 120 | 300 | 600 | 1200 | 3600 | 7200 | 14400 | 28800 | 45000
Sparse fishline | 7.58 | 1.21 | 1.15 | 1.15 | 1.14 | 1.13 | 1.09 | 1.04 | 0.95 | 0.78 | 0.65 | 0.52 | 0.37 | 0.27
Dense fishline | 8.55 | 2.69 | 2.64 | 2.64 | 2.61 | 2.58 | 2.51 | 2.4 | 2.25 | 1.93 | 1.7 | 1.47 | 1.23 | 1.05
Chart: Irradiation time [h]
| 0 | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.8 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7 | 7.75
Sparse fishline | 6.47 | 6.66 | 6.78 | 6.87 | 6.95 | 7.05 | 7.19 | 7.25 | 7.35 | 7.43 | 7.52 | 7.58 | 7.6 | 7.6 | 7.59 | 7.58
Dense fishline | 6.42 | 6.62 | 6.73 | 6.82 | 6.91 | 6.99 | 7.18 | 7.29 | 7.51 | 7.67 | 7.93 | 8.13 | 8.29 | 8.4 | 8.51 | 8.55
Gas flow
Gas inlet
Gas outlet
Kapton spacer: 3 x 3 x0.25mm
Dark current decay after 10h
@7 kHz/cm2 radiation

## Slide 13: Gas flush simulation
Mass percentage of pad spacer MRPC
Mass percentage of fishing line spacer MRPC

## Slide 14
Strip-MRPC
Inner zone-MRPC
Prototype design for CBM-TOF
Two-dimensional code of MRPC

## Slide 15
Now delayed to 2027!

## Slide 16
Maintain capability to deliver lower pass beam energies : 2.2, 4.4, 6.6,….
Enhanced capabilities in existing Halls
Increase of Luminosity 1035 - ~1039 cm-2s-1
The completion of the 12 GeV Upgrade of CEBAF was ranked the highest priority in the 2007 NSAC Long Range Plan.
New Hall
CHL-2
20 cryomodules
Add 5
cryomodules
Add 5
cryomodules
20 cryomodules
Add arc
JLab 12Gev Upgrade

## Slide 17
Full exploitation of JLab 12 GeV Upgrade
 A Large Acceptance Detector AND Can Handle High Luminosity (1037-1039)
Take advantage of latest development in detectors , data acquisitions and simulations
Reach ultimate precision for SIDIS (TMDs), PVDIS in high-x region and threshold J/y
5 highly rated experiments approved (+3)
Three SIDIS experiments, one PVDIS, one J/y production (+ three run group experiments)
Strong collaboration (250+ collaborators from 70+ institutes, 13 countries)
Overview of SoLID
The MRPC is developed for the TOF of SoLID
Main Requirements for TOF:
  /k separation up to 7GeV/c
  Time resolution < 20ps
  Rate capability > 20kHz/cm²

## Slide 18: Toward 20ps resolution: narrow gap MRPC
TOF <20 ps， the intrinsic resolution of narrow gaps MRPC is around 15ps，so the time jitter of readout electronics <13~15 ps。
Simulation indicates proper ways to design the gap thickness and arrange the stacks
Narrower gap thickness required
MRPC <20 ps，the gas gap: <0.18mm
gap number: >16

## Slide 19: Toward 20ps resolution: fast readout electronics
low

## Slide 20: Toward 20ps resolution: fast readout electronics
Fast amplifier & Waveform sampling
Electronic system time σ<7ps
Waveform sampling
Based on DRS4
Sampling freq=5 GS/s
Fast amplifier
Bandwidth=1.4GHz
 | MRPC prototype
gas gap thickness | 128 μm
number of gas gaps | 4 chambers × 8 gaps
glass material | low resistivity glass
glass thickness | 400
readout strips | 5 mm in width (2 mm clearance)
材料厚度：<0.1X0

## Slide 21: Neural network can be used to analysis time
Multilayer perceptron (MLP)
Input
Output
Activation function: g and h —— tanh
Weights:
“Dropout”: avoid overfitting
Several uniformly distributed points along the leading edge
The length of the leading edge t l
Train/validate/test set: 20/10/10 k
Tensorflow & GPU: GTX 1080 Ti
~ 10 mins for training
Time of the very first interaction:
MRPC
粒子源
固定阈值

## Slide 22
Recurrent neural networks(RNN): Long Short Term Memory network(LSTM)
Train/validate/test set: 20/10/10 k
Tensorflow & GPU: GTX 1080 Ti
F
I
G
O
f : forget gate: Whether to erase
I : Input gate, whether to write
g: gate gate, How much to write
o: output gate, How much to reveal
The length of the leading edge t l
Several uniformly distributed points along the leading edge
> 30 mins for training
Another NN：LSTM

## Slide 23: Cosmic test
Total time resolution
MRPC + fast amplifier + DRS4
Timing：CFD
Event choose：Vertical cosmic ray
E=140 KV/cm

## Slide 24: 20ps,15kHz/cm2 MRPC
detection efficiency
time resolution
detection efficiency
time resolution
rate
We are also trying with NINOs+ FPGA TDC to get good resolution!

## Slide 25
PID of STCF
Physical Targets of STCF
Parameters of STCF
the asymmetry of matter-antimatter
structure of hadrons and the nature of non-perturbative strong interactions
physics beyond the standard module
Center-of-mass energy: 2-7GeV
Luminosity: ＞0.5*1035cm-2s-1
Physical requirements of STCF-PID
π/K misidentification rate < 2% at p < 2GeV/c
PID efficiency of hadrons >97% at p < 2GeV/c
Time resolution of eTOF：<44ps

## Slide 26
Structure of STCF-eTOF
Component | Dimensions(mm) | Quantity
Honeycomb | 470×(222.4-405.8)×6 | 2
PCB | 470×(226.4-409.8)×0.8 | 5
Mylar | 470×(222.4-405.8)×0.25 | 8
Float glass | 470×(224.4-407.8)×0.3 | 28
Gas gap | 0.14 | 24
Readout strip | 470×(222.4-405.8) | 24
470mm
222.4mm
405.8mm
MRPC
读出条
425.8mm
242.4mm
每个扇区由一个梯形MRPC组成
密封型24气隙MRPC结构

## Slide 27: STCF-eTOF主要指标
性能：
1、时间分辨：<44ps
2、PID of /k: 2GeV @3
3、TOF面积：~2.362m2单面
4、电子学道数：1024
5、电子学功耗：17mW/道
6、造价估算：inexpensive
7、材料厚度：0.1X0
8、读出电子学：TOT放大器+FPGA TDC

## Slide 28: Summary
The rate capability of MRPC assembled with low resistive glass can reach 70kHz/cm2
Sealed MRPC was developed in our university:
-- Promote gas exchange in gas gap
-- Sealed MRPC (Kapton pad spacer) can solve creepage on the spacer at high
rate condition
-- Reduce gas consumption greatly
-- Has been used in CBM-TOF and CEE-eTOF
The time resolution of narrow gap MRPC with fast electronics can be better than 20ps
MRPC-TOF is a good choice for STCF-eTOF!!!

## Slide 29
Thanks for your
attention !